## Slide 1: CERN Mock-up test results
Long run at ln2
P. Barrillon(1), L. Boistay(2), F. Hubaut (1), P. Pralavorio(1), P. Skensved(3), M. Van Uffelen(1), Cryolab(2)
(1) CPPM, (2) CERN,(3) Queen’s University
Setting up the Mock-up
Overall conditions of the test
Stress test
Long stay test
Conclusion
Next step

## Slide 2
Mock-up
Introduction
 | MU_CShort | MU_CLong | MU_Warm
Purpose | Robustness against ice formation, source blocked, bend, pipe leak |  | Behaviour in bends, DS20k Length
Temp. (K) | LN2 (77) | LN2 (77), LAr (88) | Room (290)
Location | CPPM | CERN | CPPM
Pipe lgth (m) | 4 | 2 | 15 (~DS-20k)
Nb bends | 2 | 1 | 15 (11, 7)
Runs | 09+11/22, 03/23 | 1 month @ June 23 | 03-04/23
Determine robustness of calibration system at cold, and during a long period (LN2 then LAr).
Huge thanks to the CERN Cryolab for providing and running the Cryostat, the LN2 and the LAr

## Slide 3: Mock-up
Test’s schedule
LN2
Dewar1
LN2
Dewar2
Cryo
Pseudo–source
L = 5.5cm
f = 2.5 cm
M= O(100) g
In CERN Cryolab, building 159
30/05-2/06 : Installation
5-9/06 : Settings
9-26/06 : LN2 tests
26/06-03/07 : Warming up Cryo to put LAr
03-10/07 : LAr tests (TBD)

## Slide 4: Mock-up
Setup
LN2
Dewar1
4 PT100 probes (A, B, C, D)
1.3 meter of useful tube (from A to D)
The cryostat is not fully adiabatic  Need constant refill with LN2
The tube cap is not fully hermetic because of rope and gN2 flushing
Rope
gN2 flush
DS4
DS2

## Slide 5: Mock-up
RUN | Temperature of A, B, C, D | Tension | Conclusion
CPPM short run, Poorly isolated cryostat (0.16 m3) |  |  | Left gN2 flushing at 60L/h TA <= TLAr for only 5h after filling Tension 12-40N
CERN long run, Vacuum isolated cryostat (0.4 m3) |  |  | Left gN2 flushing at 60L/h TA, TB, TC always below TLAr once filled Tension 5-25N
Comparison with CPPM MOCK-UP (first 8 hours)
Motorized Systems work at CERN.
Dewar1
Dewar2

## Slide 6: Overall conditions of the test
Date | 09/06 | 10/06 | 12/06 | 13/06 | 26/06
Left Side (DS4) |  |  |  |  | 
Right Side (DS2) |  |  |  |  | 
Pictures are cropped to be visible so do not pay attention on sizes, more on shapes.
Ice formation mitigation
1- Ice formation on both sides (09/06)  Flush hot gN2 on left at 250L/h (4x nominal)
 Left side ice turns into frost
2- LN2 inside (-50cm) tube prevent gN2 to reach right side  T pipe to flush left and right (250L/h)
 Ice on right side turn into frost
Frost
Ice
LN2
The rope on the corner of tube to take the picture, in center usually
Ice turns into frost after 1 day of
large gN2 flush on both sides  Stable conditions for test
Frost
(T)
(T)

## Slide 7
Test conditions (18 days)
Temperature of PT100 and tank filling
Chart: %LN2 in tank vs time
0
50
100
150
300
200
250
gN2 flow (L/h)
Filling process is automatic (no auto-filling on the weekends and on Wednesday)
Flushing at 120L/h (2x60 with T)
Tank refills
D
A,B,C
End of cooling
Slope
0.5%/h
Prepa
Stable
TA, TB, TC always below TLAr for 18 days

## Slide 8
Stress test with LN2
Program
Cold Garage
The goal is to check robustness of the Mock-up
Back and forth (b&f) from the cold garage to PT100 A (2x 1,3m) with a ~1min stop at the edges (~1 cm/s)
35 b&f/day (~3h) for 8 days In DS20K, 8 sources so 4b&f and 20m tube  280 b&f  70x DS20k  ~740m  5x DS20k
DS4
DS2

## Slide 9
Stress test with LN2
TYPICAL RUN
Typical tension for DS2
Tension ~23N
Tension ~16N
Following motor’s tension
6N
Typical tension for DS4
Cold garage
D
In tube LN2 level
C
B
A
2 typical tension depending puller motor
DS4
DS2
DS2 tension around 24N, no visible impact of tube geometry
DS4 tension lower, and reflects the tube geometry
90s

## Slide 10
Stress test with LN2
Results (First week)
13/06
14/06
15/06
16/06
40N spike
10min break during the test
Tension ~23N
Tension ~16N
Following motor’s tension
6N
DS2 tension : two days at 25-30N at worst
25°C/40%
23,5°C/40%
23,5°C/40%
24,5°C/43%

## Slide 11
Stress test with LN2
Results (Second week)
19/06
20/06
21/06
22/06
10min break during the test
The tension is stable
Tension ~23N
Tension ~16N
Following motor’s tension
6N
24,5°C/47%
25°C/54%
26°C/53%
24°C/60%

## Slide 12
Long stay test with LN2
Program for the test
The goal is to check how the motors react when the source is left for a long time at a PT100 step.
We left the source at different spots for different periods of time (1min, 10min, 30min, 1h, 2h, a night, a weekend)
Source | 57Co | 133Ba | 22Na | 137Cs | 60Co | AmBe | AmC | AmLi
Energy (keV) | 122 | 356 | 511 | 662 | 1173 | [0.2, 12] | [2, 7] | [0, 2]
Activity (side) (kBq) | 18 | 1,9 | 0,36 | 2,2 | 0,36 | 0,14 | 0,15 | -
Activity (bottom) (kBq) | 100 | 5 | 0,67 | 4,6 | 0,6 | 0,18 | 0,18 | -
Duration of calibration (h) | 3,84 | 18,72 | 23,52 | 36 | 74,4 | 200 | 200 | -
Time on each spot | 12’ | 1h50’ | 2h20 | 3h45’ | 8h | 22h | 22h | -

## Slide 13
Long stay test with LN2
Tension after stop
When the source is moved from its staying position, extract maximal tension 10s after the start of the move
Tension plot during one test
1min
10min
30min
A
B
C
D
C
B
A

## Slide 14
Long stay test with LN2
Tension after stop
Need to be checked

## Slide 15
Long stay test with LN2
Tension after stop
Tension is not related to the duration of the stay at the same position

## Slide 16
Conclusion
LN2 Mock-up test
Liquid in tank | LN2
Speed of the source (cm/s) | 1
Position Accuracy (cm) | 
Tension (N) | 15-30
Ice formation (block) | Mitigated with gN2
Total Length for one source (m) | 800
Total nb of back&forth / pipe | 280
Liquid in tank | LN2
Speed of the source (cm/s) | 1
Position Accuracy (cm) | 
Tension (N) | 15-30
Ice formation (block) | Mitigated with gN2
Total Length for one source (m) | 800
Total nb of back&forth / pipe | 280
Run at CERN at LN2 for 18 days (1,3m of useful tube)
  Liquid in the tube 50cm from top  T pipe
  Ice formation mitigated using gN2  Frost
  Average tension 15-30N during stress test
  No impact on tension when stopping (from 1min to 2 days)
Robust behavior of calib system at LN2 temperature for 2 weeks

## Slide 17
Next test :
LAr for 5 days
Warming up the cryostat for 7 days : from LN2 temperature to above LAr.
We expect to start 3rd of July :
  No liquid in tube  gN2 60L/h
  Redo stress and calib tests
Liquid in tank | LAr
Speed of the source (cm/s) | ?
Position Accuracy (cm) | ?
Tension (N) | ?
Ice formation (block) | ?
Total Length for one source (m) | ?
Total nb of back&forth / pipe | ?