## Slide 1: LAr tests
From 03/07 to 07/07
Lucas

## Slide 2: Summary :
DS2
(right)
DS4
(left)
Date | Hour | Ice Formation | °C | Tension before any test (N) | Voltage (V) | Current (A)
Mon – 03/07 | 13:35 | None | A,B at -183,95°C (LAr) C -145 D -68 | DS2 : 5 DS4 : 5 (relaxed) | DS2 : 24 DS4 : 24 | DS2 : 0.08 DS4 : 0.09
Tue – 04/07 | 10:50 | Intended | A, B LAr Temperature C -137.37 °C D -66,51°C | DS2 : 15 DS4 : 5 (not relaxed) | // | //
Wed – 05/07 | 14:21 | Intended | A, B LAr Temperature C -130.71 °C D -63,13 °C | DS2: 20 DS4: 5 (not relaxed) | // | //
Fri – 07/07 | 14:00 | Same | A,B LAr Temperature C -121.33 °C D -60,8 °C | DS2 : 25 DS4 : 7 (not relaxed) | // | //
+2 humidity sensors

## Slide 3: Current analysis
LN2 tests
LAr tests
Motors used 3.9%
on 18 days
Motors used 3%
on 3 days
Heating
process
DS2
DS4
 Motors used ~4%
of the total time during LN2 tests

## Slide 4: Temperature changes since Friday 30/06 at 9AM
Monday 26/06 : Start the heating process. Small heater placed in the tank.
Thursday 29/06 : Heating process accelerated : larger heater inside the tank, lid of the tank removed
Friday 30/06 : End of the active heating process, letting tank heat with ambient temperature.
Monday 03/07 : Filling of the cuve with LAr
Friday 30/06 – 18h
Saturday
Sunday
Monday
LAr filling
Tuesday
Wednesday
LAr temperature
Thursday
Friday

## Slide 5: LAr in tank curve
Chart: DarkSide level during long test run in LAr
Tank level decreasing linearly
(around 4%/day)

## Slide 6: Flushing LN2 to see effect on ice :
Time | Monday 03/07 | Tuesday 04/07 | Wednesday 05/07 | Friday 07/07
Right Side (DS2) |  |  |  | 
Left Side (DS4) |  |  |  | 
T shaped gN2 pipe | NO | NO | NO | NO
gN2 Total flow | 60 L/h | 60L/h | 120L/h | 60L/h
LN2 in tube | None | None | None | None
%LN2 in tank* | 61% | 59% | 54% | ??
Pictures are cropped to be visible so do not pay attention on sizes, more on shapes.
* No refill possible this time

## Slide 7: Tension during stress test 03/07
Again repositioning
of motors
Repositioning
of motors
At beginning, tension very high  Reposition of the motors to be in straightly in front of the cap centers  Once again after adding humidity sensors
Tension value about same value as LN2 testing
 Tension value stable and comparable to LN2
Avg DS2 tension
after repositioning
21N
Avg DS4 tension
after repositioning
15N

## Slide 8: Long stay test (one night)
Source left at D for 16h, no high tension after moving it back to A (DS4)
T=15N
T=15N
Max tension 16N

## Slide 9: Humidity test 1/3
Regularly moving source from D to A to D, for different gN2 flow values
Left Humidity rising only if gN2 < 30L/h
Right humidity rise only if gN2 < 120L/h
  No difference 0 and 30 L/h
  Regular drop from 30L/h (80%) to 120L/h (~0%)
No humidity changes for gN2 flow at 120 L/h, no matter movement of pseudo-source.
gN2 flow at 0L/h
gN2 flow at 30L/h
gN2 flow at 120L/h
gN2 flow at 60L/h
A
A
A
A
A
A
D
D
D
D
D
D
D
DS4 (left)
DS2 (right)

## Slide 10: Humidity test 1/3 – Without moving source
DS4 (left)
DS2 (right)
0
30
60
120
0
30
60
120
Reset Flushing 250L/h
120
gN2 flushing (L/h)
Source position at D
Source position at A
Vary gn2 flushing F and measure humidity when source at two positions : top-left (D), bottom (A)
Determine F which remove humidity “instantly” (<=10min)  Fthreshold
Source Top-left : Vbefore small  60 < Fthreshold < 120 L/h  Measure : Fthreshold = 100L/h
Source Bottom: Vbefore larger  120 < Fthreshold < 250 L/h  Measure : Fthreshold = 150L/h
 Assuming same scaling, Fthreshold for DS20k with source at bottom might be around 300L/h
D
C
B
A
D
C
B
A
Vbefore(D)
4xVbefore(D)
Source
Source

## Slide 11: Humidity test 2/3
Pseudo-source in position A
Blow in tube from both sides to increase humidity at maximum and remove gN2 = Extreme nonrealistic scenario
After 5min, both humidity are at 100%
A D  A : +23N/+15N wrt normal
After 15min, put back the top caps and gN2 at 120L/h
After 1h30, both humidities at 0%
A D  A : +20N/+15N wrt normal but back to normal after 3 times
Humidity at 100%
Humidity at 0%
Avg DS2 tension in normal conditions
21N
Avg DS4 tension in normal conditions
15N
A
A
D
A
D
A
A
D
D
A

## Slide 12: Humidity test 3/3
Humidity at 100%
Humidity at 0%
Avg DS2 tension in normal conditions
21N
Avg DS4 tension in normal conditions
15N
A
A
D
A
D
A
A
D
A
CG*
HG*
Pseudo-source in position B
Blow in tube from both sides to increase humidity at maximum and remove gN2 = Extreme nonrealistic scenario
After 5min, both humidity are at 100%
A D  A : +12N/+11N wrt normal
After 15min, put back the top caps and gN2 at 120L/h
After 1h30, both humidities at 0%
A HG*  CG* : +12N/+9N wrt normal
*HG : Hot Garage
*CG : Cold Garage

## Slide 13: Humidity test 3/3 – After night
Avg DS2 tension in normal conditions
21N
Avg DS4 tension in normal conditions
15N
CG*
A
D
Pseudo-source in position Cold garage for 1 night
CG*  D  A : +8N/+5N wrt normal
*HG : Hot Garage
*CG : Cold Garage
 With long-time flushing, maximum tension due to ice formation decreases
Data from previous slide
18h before

## Slide 14: Endoscope pictures
DS4
DS2
After long blowing (100% humidity)
After 1h 15mn of gN2 @120 l/mn
After 2nd long blowing (100% humidity)

## Slide 15: Ice/Frost identification DS2
Frost pieces coming off
The top part is frost (~1mm thick)
Soft frost removed with the stick
Illusion of thickness
due to camera
Correct thickness

## Slide 16: Ice/Frost identification DS4
Frost pieces coming off
Frost pieces coming off
Illusion of thickness
due to camera
Correct thickness
The top part is frost (~1mm thick)